Great Bay (GRB) NERR Water Quality Metadata

April - December 2014

Latest update: February 29, 2016

I. Data Set & Research Descriptors

1. Principal investigators and contact persons

Thomas K. Gregory

Research Scientist

email: tom.gregory@unh.edu

Ocean Process Analysis Lab

University of New Hampshire

Durham, NH 03824

603-862-5136

Jonathan Pennock, Ph.D., Professor of Natural Resources, Director of the

Jackson Estuarine Laboratory and the UNH marine Program

email: [Jonathan.Pennock@unh.edu](mailto:Jonathan.Pennock@unh.edu)

Chase Ocean Eng. Room 102

U.N.H.

Durham, NH 03824

Phone: (603) 862-2921

Paul E. Stacey

Research Coordinator

Great Bay National Estuarine Research Reserve

email: Paul.Stacey@wildlife.nh.gov

New Hampshire Fish & Game Department

225 Main Street, Durham, NH 03824

Phone (603) 868-1095

2. Entry verification

Deployment data are uploaded from the YSI data logger to a Personal Computer (IBM compatible). Files are exported from EcoWatch in a comma-delimited format (.CDF) and uploaded to the CDMO where they undergo automated primary QAQC; automated depth/level corrections for changes in barometric pressure (cDepth or cLevel parameters); and become part of the CDMO’s online provisional database. All pre- and post-deployment data are removed from the file prior to upload. During primary QAQC, data are flagged if they are missing or out of sensor range. The edited file is then returned to the Reserve for secondary QAQC where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro. The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review. It allows the user to apply QAQC flags and codes to the data, remove any overlapping deployment data, append files, and export the resulting data file for upload to the CDMO. Upload after secondary QAQC results in ingestion into the database as provisional plus data, recalculation of cDepth or cLevel parameters, and finally tertiary QAQC by the CDMO and assimilation into the CDMO’s authoritative online database. Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate. For more information on QAQC flags and codes, see Sections 11 and 12. Tom Gregory is responsible for data management.

3. Research Objectives

In 2014, YSI 6600 and EXO2 datasondes were deployed in Great Bay and in the Squamscott, Oyster and Lamprey Rivers as part of the National Estuarine Research Reserves' (NERRS) System-wide Monitoring Program (SWMP). The goal is to develop and maintain temporally intensive long-term datasets of physio-chemical parameters of water quality at locations that are representative of the Great Bay Estuarine System. The Great Bay site is relatively unimpacted, while the three tidal river sites (Lamprey, Oyster and Squamscott) have large drainage basins and are impacted by both point (wastewater treatment plants) and nonpoint sources of pollution. In addition to establishing a baseline of water quality, and understanding the spatial and temporal variability of important indicators of estuarine water quality, the data is used by a number of researchers in the analysis of physical and biological processes.

4. Research Methods

Datasondes are programmed to obtain measurements of specific conductivity,

salinity, dissolved oxygen, percent saturation, pH, temperature, water level,

and turbidity every 15 minutes. Sondes used in 2014 were model 6600 or EXO2 sondes equipped with non-vented depth sensors. The instruments are deployed continuously during ice-free seasons, except for brief periods when they are removed for cleaning, maintenance and recalibration. Pre and post-deployment calibrations are performed using the diagnostics menu of the YSI Ecowatch or Kor programs and QA/QC procedures developed by NERR Research Coordinators and YSI engineers. YSI conductivity and pH standards are used for calibration. YSI turbidity standard is used to calibrate turbidity probes.

During each deployment field measurements of temperature, salinity and dissolved oxygen are recorded using a YSI 85 field meter.

A Sutron Sat-Link2 transmitter was in use at the Oyster River station and transmits data to the NOAA GOES satellite, NESDIS ID #3B03437E. The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu/get/export.cfm

The Great Bay sonde was deployed at 0.5 meters off the bottom in a PVC tube that was attached to the stem of a mushroom anchor. This site does not currently have telemetry. The sonde logged successfully during the reporting period.

The Lamprey and Squamscott River sondes were deployed inside piling mounted PVC tubes at 0.5 meters off the bottom. Both Lamprey and Squamscott River sondes were telemetered via Nexsens transmitters using cellular technology.  These telemetry systems operated successfully during most of the reporting period.

Due to shallow depths and narrow channels, the Oyster River sonde has to be deployed with the least amount of vertical expression above bottom. This was achieved by mounting the sonde inside a short PVC tube that was attached to the stem of a mushroom anchor. This allows for the sonde to be stationed in an upright position, while the anchor is less susceptible to dragging than the previous deployment method. The sonde was deployed at 0.3 meters off the bottom. It was telemetered via a Sutron GOES satellite telemetry system. The system functioned properly throughout most of the reporting period.

5. Site Location and Character

Site #1 Great Bay (GB)

Location: Central area of Great Bay proper.

The coordinates of the deployment are 43 04' 20" N latitude and 70 52' 10" W

longitude.

Salinity range: 5-32 ppt (seasonally); 0-5ppt from high to low tide.

Temperature range: -1C to 24C (seasonally); 0-3 ( from high to low tide)

Depth: 6.5 meters at MLW

Tidal height of 2.7 meters

Bottom type: Mud and rock channel bottom

Tidal velocity: maximum 50 cm/sec

Watersheds: Squamscott, Lamprey and Winnicut Rivers plus smaller

streams. High tide influence from Little Bay and associated rivers

Pollutant influence: clean reference site

Site #2 Squamscott River (SQ)

Location: Mid channel of the Squamscott River at the Boston and Maine Railroad

bridge, Stratham, NH.

Coordinates are 43 02' 30" N latitude and 7055' 20" W longitude

Salinity range: 0-30ppt (seasonally); 5-20 ppt from high to low tide.

Temperature range: -1 C to 27 C (seasonally);difference of 0-5( from high to

low tide)

Depth: 3.5 meters at MLW

Tidal height of 2.7 m

Bottom type: Mud/oyster channel bottom

Tidal velocity: maximum 50 cm/sec

Watersheds: Exeter River, adjacent marshes

Pollutant influence: Urban stormwater, agriculture, two municipal

wastewater treatment plants, residential septic systems

Site #3 Lamprey River (LR)

Location: West bank of the tidal portion of the Lamprey River, approximately 300 m downstream of the dam at Route 108 in Newmarket, NH.

Coordinates are 43 04' 48" N latitude and 70 56' 04" W longitude.

Salinity range 0 - 27 ppt (seasonally); difference of up to 15 ppt between high and low tides.

Temperature range -1C to 27C( (seasonally); difference of up to 5C between

high and low tides.

Depth: 3.5 m

Tidal height 2.7 m

Bottom type: mud/rock

Tidal velocity: maximum 40 cm/sec

Watershed: Lamprey River

Pollutant influence: Urban stormwater, adjacent marina, upstream and downstream wastewater treatment plants, upstream agriculture

Site #4 Oyster River (OR)

Location: in the center channel of the tidal portion of the Oyster River,

approximately 300 m downstream of the head of tide dam adjacent to Jackson’s

Landing in Durham, NH.

Coordinates are 43° 08' 2.4" N latitude and 70° 54' 39.6" W longitude

Salinity range 0 –32ppt (seasonally); difference of up to 15 ppt between high and low tides

Temperature range -1C to 27C( (seasonally); difference of up to 5°C between high and low tides

Depth: 0.3 m at mlw, 3 m at highest high tides

Tidal height 2.7 m (maximum)

Bottom type: mud

Tidal velocity: maximum 40 cm/sec

Watershed: Oyster River

Pollutant influence: urban stormwater, mooring field and crew dock, downstream wastewater treatment plant, upstream agriculture, residential on-site sewage disposal.

6. Data collection period

Great Bay data collection began April 21, 2014 at 1230.

Squamscott River data collection began April 21, 2014 at 1300.

Lamprey River data collection began April 21, 2014 at 1000.

Oyster River data collection began April 18, 2014 at 1515.

The instruments are removed from the water during the winter months due to non-navigable conditions caused by ice and the removal of channel markers. Icing is particularly severe in the rivers and is harmful to instruments, boats and telemetry equipment.

DEPLOYMENT DATES 2014

GREAT BAY

Deploy date and time Retrieve date and time

4/21/14 1230 5/22/14 1300

5/22/14 1330 6/17/14 1115

6/17/14 1130 7/11/14 1130

7/11/14 1145 8/4/14 1200

8/4/14 1230 9/5/14 0900

9/5/14 0930 10/9/14 1245

10/9/14 1300 11/6/14 1215

11/6/14 1230 12/5/14 1300

LAMPREY RIVER

Deploy date and time Retrieve date and time

4/21/14 1000 5/23/14 1015

5/23/14 1030 6/19/14 0900

6/19/14 0915 7/9/14 0930

7/9/14 0945 7/29/14 1400

7/29/14 1415 9/12/14 1400

9/12/14 1415 10/20/14 1315

10/20/14 1330 11/11/14 1430

11/11/14 1445 12/5/14 1245

OYSTER RIVER

Deploy date and time Retrieve date and time

4/18/14 1515 5/21/14 0930

5/21/14 0945 6/19/14 0645

6/19/14 0700 7/14/14 1315

7/14/14 1330 8/22/14 1115

8/22/14 1145 9/24/14 1100

9/24/14 1130 10/12/14 1215

10/17/14 1030 11/08/14 2245

11/10/14 1415 12/5/11 0400

SQUAMSCOTT RIVER

Deploy date and time Retrieve date and time

4/21/14 1300 5/6/14 1100

5/6/14 1115 6/9/14 1300

6/9/14 1315 7/17/14 1245

7/17/14 1300 8/25/14 1130

8/25/14 1145 9/26/14 1000

9/26/14 1030 10/21/14 1045

10/21/14 1100 11/11/14 1245

11/11/14 1300 12/5/14 1215

7. Distribution:

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: <http://www.nerrsdata.org/>; *accessed* 12 October 2012.

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page [http://cdmo.baruch.sc.edu/](http://cfcdmo.baruch.sc.edu/).  Data are available in comma delimited format.

8. Associated researchers and projects:

NOAA's National Estuarine Research Reserve System (NERRS)

NERRS System-wide Monitoring Program (SWMP)

Evaluation of remote data acquisition technologies and advanced water quality sensors. Dr. Richard Langan, CICEET director, and Jeremy LeClair, U.N.H. Funded by CICEET.

Advanced optical monitoring technologies. Ru Morrison. Funded by Center for Ocean Observing and Analysis, U.N.H.

Smelt spawning studies. Kathy Mills, NH Fish & Game, and David Berlinsky, U.N.H. Funded by by NOAA Office of Protected Resources via a sub-contract from Maine Department of Marine Resources

Comprehensive water quality, organic nitrogen and photosynthetically active radiation (PAR) monitoring studies – Dr. Jonathan Pennock, Jackson Estuarine Laboratory. Supported by the New Hampshire Estuaries Project

Eelgrass modeling studies - Dr. Fred Short, Jackson Estuarine Laboratory. Supported by the New Hampshire Estuaries Project and the New Hampshire Port Authority.

Bathymetric modeling and tidal elevation studies conducted by the NOAA – Dr. Larry Mayer, UNH Center for Coastal Ocean Mapping. Supported by the UNH-NOAA Joint Hydrographic Center.

Oyster reef mapping and restoration – Dr. Ray Grizzle, Jackson Estuarine Laboratory. Supported by NH Fish and Game, the NOAA-UNH Joint Hydrographic Center and the Center for Coastal and Ocean Mapping.

Microbial source tracking studies using ribotyping - Dr. Stephen Jones, Jackson Estuarine Laboratory. Supported by NH DES, NHEP and CICEET

Lobster migration and behavior research - Dr. Winsor Watson and Dr. Hunting Howell, UNH Zoology Department. Ten years of studies supported by USDA and Sea Grant that track lobster abundance, movement and behavior in relation to physical and biological variables in the Great Bay Estuary.

EPA national Coastal Assessment Program - Dr. Stephen H. Jones, Jackson Estuarine Laboratory. Funded by the US-EPA.

Anadramous and juvenile fish population assessments – Cheri Patterson, NH Fish and Game Department and Great Bay NERRS. Supported by NH Fish and Game.

II. Physical Structure Descriptors

1. Sensor specifications

YSI 6600 model sondes with ROX DO sensors were deployed at the GB site for 2014. YSI EXO2 sondes were deployed at LR and SQ for 2014. A 6600 model sonde was used for the 4/18/2014 deployment at OR and EXO2 sondes were used for the rest of the year.

YSI 6600 sonde

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model #: 6560

Range: -5 to 45 °C

Accuracy: +/-0.15 °C

Resolution: 0.01 °C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode cell with autoranging

Model #: 6560

Range: 0 to 100 mS/cm

Accuracy: +/-0.5% of reading + 0.001 mS/cm

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependent)

Parameter: Salinity

Units: parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: +/- 1.0% of reading or 0.1 ppt, whichever is greater

Resolution: 0.01 ppt

Parameter: Dissolved Oxygen % saturation

Units: percent air saturation (%)

Sensor Type: Optical probe w/mechanical cleaning

Model # 6150 ROX

Dissolved Oxygen, % Specifications

Range: 0 to 500%

Resolution: 0.1%

Accuracy:

0 to 200%: ±1% of reading or 1% air saturation, whichever is greater

200 to 500%: ±15% of reading

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature and salinity)

Units: milligrams per Liter (mg/L)

Sensor Type: Optical probe with mechanical cleaning

Model # 6150 ROX

Range: 0 to 50 mg/L

Resolution: 0.01 mg/L

Accuracy: 0 to 20 mg/L: ±0.1 mg/L or 1% of reading, whichever is greater

20 to 50 mg/L: ±15% of reading

Parameter: Non-Vented Level – Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 30 ft (9.1 m)

Accuracy: +/- 0.06 ft (0.018 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH

Bulb probe

Units: pH units

Sensor Type: Glass combination electrode

Model #: 6561 or 6561FG

Range: 0 to 14 units

Accuracy: +/- 0.2 units

Resolution: 0.01 units

Parameter: Turbidity

Units: nephelometric turbidity units (NTU)

Sensor Type: Optical, 90 ° scatter, with mechanical cleaning

Model #: 6136 or 6026

Range: 0 to 1000 NTU

Accuracy: +/- 5 % reading or 0.3 NTU (whichever is greater)

Resolution: 0.1 NTU

YSI EXO2 Sonde:

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model#: 599870-01

Range: -5 to 50 C

Accuracy: -5 to 35: +/- 0.01, 35 to 50: +/- .005

Resolution: 0.01 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode cell with autoranging

Model#: 599870-01

Range: 0 to 200 mS/cm

Accuracy: 0 to 100: +/- 0.5% of reading or 0.001 mS/cm; 100 to 200: +/- 1% of reading

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependant)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 psu

Accuracy: +/- 1.0% of reading pr 0.1 ppt, whichever is greater

Resolution: 0.01 psu

Parameter: Dissolved Oxygen % saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/- 1% of the reading or 1% air saturation, whichever is greater 200-500% air saturation: +/- 5% or reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/-0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +/- 5% of the reading

Resolution: 0.01 mg/L

Parameter: Non-vented Level - Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 33 ft (10 m)

Accuracy: +/- 0.013 ft (0.04 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH

Units: pH units

Sensor Type: Glass combination electrode

Model#: 599701(guarded) or 599702(wiped)

Range: 0 to 14 units

Accuracy: +/- 0.01 units within +/- 10° of calibration temperature, +/- 0.02 units for entire temperature range

Resolution: 0.01 units

Parameter: Turbidity

Units: formazin nephelometric units (FNU)

Sensor Type: Optical, 90 degree scatter

Model#: 599101-01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +/-2% of reading (whichever is greater); 1000 to 4000 FNU +/-5% of reading

Resolution: 0 to 999 FNU: 0.01 FNU, 1000 to 4000 FNU: 0.1 FNU

**Depth Qualifier:**

The NERR System-Wide Monitoring Program utilizes YSI data sondes that can be equipped with either depth or water level sensors.  Both sensors measure water depth, but by convention, level sensors refer to atmospherically vented measurements and depth refers to non-vented measurements.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.03 cm for every 1 millibar change in atmospheric pressure, and is eliminated for level sensors because they are vented to the atmosphere throughout the deployment time interval.

Beginning in 2006, NERR SWMP standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the SWMP calibration sheet or Digital Calibration Log.  This offset procedure standardizes each depth calibration for the entire NERR System.  If accurate atmospheric pressure data are available, non-vented sensor depth measurements at any NERR site can be corrected.

In 2010, the CDMO began automatically correcting depth/level data for changes in barometric pressure as measured by the Reserve’s associated meteorological station during data ingestion. These corrected depth/level data are reported as cDepth and cLevel, and are assigned QAQC flags and codes based on QAQC protocols. Please see sections 11 and 12 for QAQC flag and code definitions.

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu). These units are essentially the same and for SWMP purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the NERR System will assign psu salinity units for all data regardless of sonde type.

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for SWMP purposes they will be considered equivalent. Moving forward, the NERR System will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

1. Coded variable definitions:

Sampling station Site code Station code

Central Great Bay GB grbgbwq

Squamscott River SQ grbsqwq

Lamprey River LR grblrwq

Oyster River OR grborwq

11. QAQC flag definitions –

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F\_). During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range. All remaining data are then flagged 0, passing initial QAQC checks. During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5 Outside High Sensor Range

-4 Outside Low Sensor Range

-3 Data Rejected due to QAQC

-2 Missing Data

-1 Optional SWMP Supported Parameter

0 Data Passed Initial QAQC Checks

1 Suspect Data

2 *Open - reserved for later flag*

3 Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

4 Historical Data: Pre-Auto QAQC

5 Corrected Data

12. **) QAQC code definitions** – This section details the secondary QAQC Code definitions used in combination with the flags above. Include the following excerpt:

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column. There are three (3) different code categories, general, sensor, and comment. General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data. Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an \* below) can be applied to the entire record in the F\_Record column.

General Errors

GIC No instrument deployed due to ice

GIM Instrument malfunction

GIT Instrument recording error; recovered telemetry data

GMC No instrument deployed due to maintenance/calibration

GNF Deployment tube clogged / no flow

GOW Out of water event

GPF Power failure / low battery

GQR Data rejected due to QA/QC checks

GSM See metadata

Corrected Depth/Level Data Codes

GCC Calculated with data that were corrected during QA/QC

GCM Calculated value could not be determined due to missing data

GCR Calculated value could not be determined due to rejected data

GCS Calculated value suspect due to questionable data

GCU Calculated value could not be determined due to unavailable data

Sensor Errors

SBO Blocked optic

SCF Conductivity sensor failure

SDF Depth port frozen

SDG Suspect due to sensor diagnostics

SDO DO suspect

SDP DO membrane puncture

SIC Incorrect calibration / contaminated standard

SNV Negative value

SOW Sensor out of water

SPC Post calibration out of range

SQR Data rejected due to QAQC checks

SSD Sensor drift

SSM Sensor malfunction

SSR Sensor removed / not deployed

STF Catastrophic temperature sensor failure

STS Turbidity spike

SWM Wiper malfunction / loss

Comments

CAB\* Algal bloom

CAF Acceptable calibration/accuracy error of sensor

CAP Depth sensor in water, affected by atmospheric pressure

CBF Biofouling

CCU Cause unknown

CDA\* DO hypoxia (<3 mg/L)

CDB\* Disturbed bottom

CDF Data appear to fit conditions

CFK\* Fish kill

CIP Surface ice present at sample station

CLT\* Low tide

CMC\* In field maintenance/cleaning

CMD\* Mud in probe guard

CND New deployment begins

CRE\* Significant rain event

CSM\* See metadata

CTS Turbidity spike

CVT\* Possible vandalism/tampering

CWD\* Data collected at wrong depth

CWE\* Significant weather event

13. Post deployment information : Included in annual metadata document.

Great Bay

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Deploy Date | SpCond | ROXDO1 | pH7 | Turb |
| 4/21/2014 | 50.07 | 102.4 | 7.25 | -0.2(0.0) |
| 5/22/2014 | 50.03 | 101.9 | 6.75 | -0.1(0.0) |
| 6/17/2014 | 51.65 | 100.2 | 7.11 | 0.3(0.0) |
| 7/11/2014 |  | 104.5 | 7.48 | 0.1(0.0) |
| 8/4/2014 | 49.1 | 95.7 | 7.17 | 0.2(0.0) |
| 9/5/2014 | 50.06 | 101.2 | 6.88 | 0.1(0.0) |
| 10/9/2014 | 49.01 | 106 | 6.99 | 0.2(0.0) |
| 11/6/2014 | 49.86 | 99.3 | 7.04 | 0.2(0.0) |

Lamprey River

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Deploy Date | SpCond | ROXDO1 | pH7 | Turb |
| 4/21/2014 | 49.45 | 95.8 | 7.12 | -0.45(0.0) |
| 5/23/2014 | 50.16 | 103.9 | 7.13 | 0.58(0.0) |
| 6/19/2014 | 52.45 | 98.6 | 7.01 | -0.35(0.0) |
| 7/9/2014 | 0 | 101.7 | 7.08 | -0.42(0.0) |
| 7/29/2014 | 49.04 | 102.2 | 7.07 | -0.04(0.0) |
| 9/12/2014 | 48.06 | 98.7 | 7.08 | -0.82(0.0) |
| 10/20/2014 | 51.97 | 103 | 7.15 | 0.0(0.0) |
| 11/11/2014 | 49.85 | 98.1 | 7.17 | 0.08(0.0) |

Oyster River

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Deploy Date | SpCond | ROXDO1 | pH7 | Turb |
| 4/18/2014 | 50.95 | 98.2 | 7.32 | 0.0(0.0) |
| 5/21/2014 | 47.4 | 105.2 | 7.14 | 0.34(0.0) |
| 6/19/2014 | 49.79 | 102.7 | 7 | -0.38(0.0) |
| 7/14/2014 | 50.17 | 101 | 7.11 | 0.07(0.0) |
| 8/22/2014 | 49.8 | 99.7 | 7.05 | 0.25(0.0) |
| 9/24/2014 | 48.69 | 100.7 | 7.01 | 0.05(0.0) |
| 10/17/2014 | 50.78 | 105 | 7.02 | -0.3(0.0) |
| 11/10/2014 | 50.36 | 103.3 | 7.04 | 0.22(0.0) |

Squamscott River

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Deploy Date | SpCond | ROXDO1 | pH7 | Turb |
| 4/21/2014 | 50.03 | 106.5 | 7.09 | 0.01(0.0) |
| 5/6/2014 | 49.492 | 100.2 | 7.17 | 0.54(0.0) |
| 6/6/2014 | 49.18 | 103.7 | 7.18 | -0.07(0.0) |
| 7/17/2014 | 51.7 | 98.2 | 7.12 | 0.6(0.0) |
| 8/25/2014 | 49.45 | 97.2 | 7.11 | 0.2(0.0) |
| 9/26/2014 | 49.8 | 99.9 | 7.13 | -0.27(0.0) |
| 10/21/2014 | 51.19 | 104.5 | 7.01 | 0.04(0.0) |
| 11/11/2014 | 49.99 | 100.8 | 7.1 | 0.05(0.0) |

14. Other remarks/notes

a) Slight shifts in data are sometimes correlated with sonde exchanges. These shifts are most noticeable in pH, specific conductivity and salinity, and may be related to sensor drift (e.g., due to fouling).

b) Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform. Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings. If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

c) Unusually high dissolved oxygen values can be the result of a plankton bloom brought on by a large volume of fresh water entering the system during and after heavy rains.

1. Heavy rain can decrease salinity in the rivers to near fresh water levels. Local rainfall data can be found at <http://www.weather.unh.edu/>
2. Only turbidity data beyond the upper range of the probe was rejected in the initial QA/QC. However, due to frequent unreliability of YSI turbidity probes, and based upon historical record, any recorded value over 300 NTU should be considered suspect or false.

Remarks on [GSM] coded data:

Oyster River 6/19 deployment: The deployment starting 6/19 7:00 and ending 7/14 13:30 is comprised of three raw files, although it is from the same sonde and the same deployment. The sonde was stopped two times during the 6/19 deployment to work on telemetry issues, data at 6/20 9:45 and 7/8 11:30 represent the sonde being started back up. When looking at the raw data for this deployment there will be three raw files with start dates of 6/19, 6/20 and 7/8.

GRB data from 7/11 11:45 to 8/4 12:00 is marked either 1 SSM or -3 STF. The temp sensor ended up failing during this deployment however it wasn’t a sudden failure and it seems that portions of the deployment have reasonable data.

GRB SpCond and salinity data from 8/27 4:45 to 8/28 23:00 is odd. This may be due to biological activity, specially sea squirts that are typical at this site.

GRB Depth from 10/9 10:00-30 is odd. The sonde did not descend to the correct depth at first due to unknown reasons. However, after 10:30 depth returns to normal.